The End of Patient Shielding?

An Update on Gonad and Fetal Shielding Recommendations

Author: Kelli Welch Haynes, EdD, RT(R), FASRT, FAEIRS

Outline

- Introduction: AAPM Position Statement & NCRP Statement No. 13

- Rationale and Scientific Support for Shift in Guidance

    - Tube design, collimators, AEC

    - Tissue weighting factors

    - Internal scatter

    - Decreased dose with digital

    - Negative impact of shielding on image quality

- Response from Professional Organizations

  - State regulations

  - National regulations

  - Federal regulations

  - Accreditation changes

  - ARRT removal of shielding (2022)

  - AAPM CARES Modules

- Impact of updated recommendation on technologists’ practice

  - Explaining and educating patients on the new recommendations

  - Technical considerations for enhancing patient protection independent of shielding

- Conclusion

Word count: XX

Credits: 0.5

Learning Objectives:

After completing this article, learners should be able to:

- Differentiate the position statements on gonadal and fetal shielding

- Explain the scientific rationale behind the elimination of routine shielding in diagnostic radiography

- Educate patients on the shielding recommendations

Highlights

- In 2019, the American Association of Physicists in Medicine released a position statement stating gonadal and fetal shielding during diagnostic imaging should be discontinued, on the basis that it offers little to no benefit and may be harmful under particular conditions. - This guidance is supported by evidence that the gonads are not as radiosensitive as once believed and that shielding may conceal pathologic findings, may hinder the imaging system's automatic exposure control, and does not decrease scatter within the patient’s body. - The American Society of Radiologic Technologists Board of Directors released an independent statement that generally indicates the risks of shielding outweigh the benefits; however, it supports gonadal and fetal shielding when it is safe and appropriate, when it reduces excess radiation exposure, and when it increases patient comfort and confidence. - Radiologic technologists need to navigate guidance in practice, adopt strategies to protect patients from radiation exposure, and educate patients on the current practice for gonadal and fetal shielding.

Introduction

Since the 1950s, radiologic technologists have routinely used lead shields to protect patients from radiation exposure and to reduce the risk of passing radiation-induced mutations on to future generations. Shielding patients is one of the “cardinal rules” of radiation protection, or time, distance, and shielding: minimize time, maximize distance, and shield the patient. However, recent scientific evidence has prompted the radiologic science community to reassess the benefits of protective shielding. In April 2019, the American Association of Physicists in Medicine (AAPM) published the Position Statement on the Use of Patient Gonadal and Fetal Shielding and shortly thereafter the article, Patient Shielding in Diagnostic Imaging: Discontinuing a Legacy Practice., The AAPM position statement stated that gonadal and fetal shielding during diagnostic imaging should be discontinued.1 This statement and article sent a shockwave through the radiologic science community. Today, we still struggle with the implications of these statements and how they impact radiography practice.

Shortly after the AAPM position statement was published, Marsh and Silosky determined that even though patient shielding is routine practice in diagnostic imaging, research demonstrates that it offers little to no benefit and may actually increase patient dose and compromise the diagnostic value of the examination.2 The basis for these conclusions comes from studies that found shielding (1) may conceal pathologic findings or hinder the imaging system's automatic exposure control (AEC), (2) may not be effective based on differences in the standard anatomical locations of the gonads, and (3) does not decrease gonadal dose caused by internal scatter radiation.1,2 The AAPM position statement has subsequently been supported by National Council on Radiation Protection and Measurements (NCRP) Statement No. 13, which recommends stopping gonadal shielding use due to technical advancements in imaging equipment and scientific evidence.

Shielding Change Rationale

The US Code of Federal Regulations first introduced patient shielding in 1976 as a concern for hereditary risks from radiation exposure, and the regulation only addressed gonadal shielding.2, At that time, it was established that fertility was not affected by radiation exposure from diagnostic examinations because the dose was too low.2 For decades, fetal and gonadal shielding in diagnostic imaging has been considered as part of the ALARA (as low as reasonably achievable) principle and, therefore, best practice. For example, Fauber found that shielding reduced male gonadal dose and promoted its application as part of everyday practice. However, Fauber’s investigation was accomplished in a controlled environment, and differences in recorded doses of radiation were negligible. Considering the development of technology in digital radiography and the recent indication of exposure risks to radiation, the NCRP and AAPM have provided guidance and their rationale for discontinuing gonadal and fetal shielding for abdominal and pelvic radiography procedures.1,3

Scientific Support

In the late 1950s, the linear no-threshold (LNT) dose-response relationship became the model used in radiation protection to estimate health effects.5, This relationship states that any dose produces a response, regardless of the size of the dose.7, However, there have not been any epidemiological studies on low-dose responses in over 15 years; all studies have been based on critical exposures at moderate-to-high levels of radiation doses. The LNT model does not consider technological advances, such as digital radiography equipment (improved detectors and generators) or the time between ionizing radiation imaging procedures. In 1959, the approximate entrance skin dose to a 4-year-old male patient having a pelvis examination was 2.5 mGy.9 By 2012, the approximate dose was 0.01 mGy. The approximate dose of a newborn kidney, ureter, bladder x-ray in 1951 was approximately 1.4 mGy and in 2020 approximately 0.07 mGy.

Human beings are continuously exposed to many sources of ionizing radiation.9 These sources can be divided into 2 main categories: natural environmental radiation and artificial (or man-made) radiation. Natural environmental radiation includes cosmic rays, terrestrial radiation, internally deposited radionuclides, and radon. Radon is the largest source of natural environmental radiation. Diagnostic x-rays constitute the largest source of man-made radiation. In 2006, the annual individual effective dose from diagnostic and interventional medical procedures was estimated as 2.9 mSv and in 2016 it was 2.3 mSv—a decrease of approximately 20% over 10 years.9 The reduction in the annual effective dose has decreased, even though the use of medical imaging has drastically increased. The Figure demonstrates the effective dose in the United States from different imaging modalities for 2016.9

- United States annual individual 
(per capita) effective dose from diagnostic and interventional 
medical procedures was estimated 
as 2.9 mSv in 2006 and 2.3 mSv in 2016—a decrease of approximately 20% over 10 years

- United States annual individual 
(per capita) effective dose from diagnostic and interventional 
medical procedures was estimated 
as 2.9 mSv in 2006 and 2.3 mSv in 2016—a decrease of approximately 20% over 10 years

In radiation therapy, techniques such as priming doses and dose fractionation treatment delivery methods have been shown to be effective in treating cancer and lowering radiation exposure to patients.8, Findings have shown that low-dose radiation exposures can accelerate immune system response and potentially prevent tumor growth and development. Priming a tumor with low doses of radiation has demonstrated an increased response to higher therapeutic doses of radiation during treatment. Dose fractionation allows for the delivery of radiation doses in multiple treatments over a longer time to amplify the obliteration of cancerous cells while reducing the injury to healthy tissues, allowing them time to recover and repair between treatments.9,

Another factor to consider is the reproductive organs' radiosensitivity compared with that of other abdominal organs. In 2007, the International Commission on Radiological Protection (ICRP) decreased the tissue weighting factor of the gonads from 0.2 to 0.08.10 However, the tissue weighting factors of the bone marrow, stomach, and colon stayed the same at 0.12, signifying these organs are more radiosensitive than the gonads.12 This information could raise the question of why we are focused on shielding a less radiosensitive organ (gonads) and not attempting to reduce radiation to more radiosensitive organs (bone marrow, stomach, and colon).11,

Shielding and Patients’ Health

In diagnostic imaging, radiation doses are not correlated with considerable damage to the fetus or gonads.1 The primary worry with radiation exposure to the reproductive organs has been heightened genetic effects (modifications in germ cells that might influence future generations).2 However, according to the ICRP, “no human studies provide direct evidence of a radiation-associated excess of heritable disease.”, Likewise, the American College of Obstetricians and Gynecologists guidelines, with support from the American College of Radiography (ACR), determined that radiation exposure through imaging—excluding fluoroscopy, computed tomography, and nuclear medicine—is at a dose much lower than the exposure associated with fetal harm.12,15

Shielding the patient does not decrease scatter created within the patient.1,13 In diagnostic imaging, the principal source of exposure to internal organs that are not within the field of view (FOV) is almost entirely from internal scatter, and placing shields on the surface of the body in an attempt to cover these organs does not impact scatter within the patient.1,14,15 Therefore, shielding the patient beyond the FOV does not protect the patient.2 Because of the centralized and surface-level location of the male gonads, shielding can protect from some of the primary beam; however, the bulk of the dose still comes from scatter generated within the patient.7,14 For female patients, because of the variations in location and depth of the ovaries from one patient to the next, gonadal shielding is often well-intended but ineffective and almost impossible to achieve without obscuring pertinent anatomy.2,13, Furthermore, as the space between the gonads and contact shield increases, the effectiveness of the shield over the ovaries decreases.14

Shielding Can Affect the Usefulness of the Examination

If the shield covers the anatomy of interest, diagnostic information may be compromised and/or a repeat examination may be needed.1 When a lead shield is positioned within the FOV, it can cover essential anatomy or pathology or create artifacts. In these cases, if the image is not repeated, the radiologist may not have crucial diagnostic information needed for the interpretation; if the image is repeated, another preventable radiation exposure event will occur.14 Evidence demonstrates this is more common than normally thought.16, Frantzen et al discovered that shields were inaccurately positioned in 91% of female pelvic examinations and 66% of male pelvic examinations.17 Furthermore, it was determined that gonadal shields were inaccurately positioned in 49% of anteroposterior pelvic examinations and 63% of frog lateral pelvic examinations, causing bony landmarks to be obscured. This finding occurred more frequently in female patients than males (76% vs 51%). When a lead shield obscures anatomy, either the technologist must repeat the examination without the shield, which causes an increase in dose, especially to those more radiosensitive organs mentioned previously, or the interpreting physician is obligated to fill in the missing information with information from prior images, which can be detrimental to the patient.2

Shielding can also have a detrimental effect on image quality due to AEC.1, Today, most x-ray imaging systems use AEC, and the occurrence of a lead apron shielding in the FOV can significantly increase x-ray output, intensifying exposure to the patient and destroying the quality of the image.18 AEC is typically used for patients aged older than 5 years to provide a more patient-specific radiation dose, theoretically decreasing the overall patient dose. However, suppose the shadow of a gonadal shield covers any part of the AEC sensor. In that case, the machine will extend the time the beam is on, which increases the exposure to organs that are located outside the shielded area, such as the bone marrow, colon, liver, stomach, and female reproductive organs.14,18 A 2020 study by Kaplan et al indicated a rise in radiation exposure to all organs of approximately 25% when using a gonadal shield with AEC equipment and a 62% increase when the lead apron was correctly placed on a female patient.18

Response From Professional Organizations

After the AAPM published their statement advocating for the discontinuation of gonadal and fetal shielding during diagnostic examinations that use radiation, the announcement was supported by various organizations, including the International Society of Radiographers and Radiological Technologists, the ACR, the Health Physics Society, and the Radiological Society of North America.3,

However, the American Society of Radiologic Technologists (ASRT) did not endorse the AAPM position statement, and the Board of Directors of the ASRT published an independent statement.19 The Board of Directors determined that shielding the gonads and the fetus during pelvic and abdominal imaging procedures may cause the risk of a repeat exposure and that the risk outweighs any possible benefit; therefore, shielding the gonads and fetus during abdominopelvic imaging should be discontinued. However, the ASRT supports gonadal and fetal shielding for procedures other than abdominaland pelvic imaging when shielding is safe and appropriate, when it reduces excess radiation exposure, and when it increases patient comfort and confidence. Ultimately, the ASRT endorses the continual use of shielding during imaging procedures where it is suitable to shield and associated with minimizing exposure to the patient. The ASRT has formed a Patient Shielding Task Force to inform all stakeholders on the justification and application of this drastic change in routine practice.

In addition, the American Registry of Radiologic Technologists (ARRT) released a statement with 4 main points.3, The ARRT’s first position statement is that health care facilities may implement varying degrees of shielding policies. Next, educational programs should continue to educate students about shielding and the effects of fetal and gonadal radiation exposure and that students and radiologic technologists have the ethical and professional duty to act in the best interest of every patient. Lastly, the ARRT’s certification examinations will not expressly refer to fetal and gonadal shielding practices but will continue to include the subjects of radiation safety and protection of patients.20

Radiography programs accredited by regional or programmatic accreditation will have minimal changes.19, The Joint Review Committee on Education in Radiologic Technology (JRCERT), the organization that accredits educational programs in radiography, does not have a specific gonadal shielding standard. However, the JRCERT released a position statement regarding patient gonadal shielding in the clinical setting. The statement indicates that, in alignment with the Standards for an Accredited Educational Program in Radiography, accredited programs must have policies in place to assure students are well-informed of the appropriate use of gonadal and fetal shielding and the best use of radiation to advocate for the health and safety of students, patients, and the general public.21

Evolving Regulatory Changes

Removing patient shielding from all federal and state regulations will take time. In response to the AAPM’s declaration in April of 2019, the US Food and Drug Administration recommended revising its regulations to revoke Code 21 of the Code of Federal Regulations.5 This involves the elimination of the endorsement that lead shielding should be used to shield the gonads during pelvic and abdominal imaging; however, this measure has been delayed due to the COVID-19 pandemic. On a state level, the NCRP recommends that state regulations be revised to follow NCRP endorsements for routine shielding of patients during pelvic and abdominal imaging.3 This is supported by the Conference of Radiation Control Program Directors, Inc., which publishes “Suggested State Regulations for Radiation Control Programs” for each state to endorse and to promote consistency of radiation control rules and regulations.

AAPM CARES Committee

As a result of dialogue from the radiologic technology community, including technologists, educators, professional organizations, and textbook authors, the AAPM created the Communicating Advances in Radiation Education for Shielding (CARES) committee to work together on the effect and inadvertent results of stopping routine shielding during diagnostic imaging examinations.4,20 In doing so, the CARES committee brought together professionals from various areas of the radiologic science profession to encourage discussions, collect varied viewpoints, examine possible changes to clinical practice, and create templates for policies and procedures to be used in diverse clinical facilities.21 Today, the CARES committee involves individuals from more than 14 professional associations worldwide, representing radiologic technologists, physicists, radiologists, and associations that direct educational programs for radiologists, radiologic technologists, and state regulators.4

Technologists’ Perspective

The notion of eliminating patient shielding is a significant shift in clinical practice, and execution will take time and education to all stakeholders.6 At facilities nationwide, policies and procedures will need to be revised to reflect current practice. The CARES committee developed a set of modules to provide information and education for technologists and other health care professionals regarding the changes to shielding policies, the radiologic technologist's role in managing safety and image quality, and how to effectively communicate the shielding changes with patients and parents/guardians.4,6

During the course of a day, radiologic technologists have numerous patient interactions, and this will be a new concept for patients as well.2 Communication is a critical component when carrying out the changes in shielding policies. Suppose a patient requests a lead apron. In that case, the radiologic technologist should answer the patient’s questions and explain the benefits and risks of shielding. Shielding may be acceptable when it does not hinder the reason for the examination; however, consistency is essential so that patients are not confused. Suppose a patient will not consent to the examination unless a shield is provided. In that case, shielding should adhere to institutional policies that reduce or remove the detrimental effect on the diagnostic potential of the examination.

There are ways the radiologic technologist can protect the patient more effectively than by providing gonadal shielding.6 By decreasing the production of internal scatter and protecting unnecessary anatomy from exposure, radiologic technologists can exponentially reduce the patient's risk. Radiologic technologists should perform the procedure, center to the anatomy of interest and the image receptor, perform proper collimation, select optimum exposure factors, minimize artifacts, and perform appropriate image processing and analysis.13 Again, the dose outside the FOV is negligible, so if technologists can ensure through proper centering, collimation, and technique that unnecessary anatomy is outside the FOV, the radiation risk to the patient drops considerably, not only from exposure to the primary beam but also the amount of internal scatter being produced. Also, in the last 50 years, 3 developments have drastically decreased patient exposure during imaging examinations: advancements in image receptors, more efficient x-ray generators; and increased filtration of the x-ray beam.3 Due to these developments, a smaller dose of radiation is needed to produce a quality image today.

Even though the NCRP has recommended discontinuing patient gonadal and fetal shielding, radiologic technologists must continue to protect themselves from radiation exposure by using protective apparel.17, Since scatter radiation from the patient is the most significant work-related hazard in diagnostic imaging, any measure that reduces scattered radiation will drastically reduce the occupational exposure of radiologic technologists.22

The best way for radiographers and radiologists to conscientiously employ the ALARA principle is to reduce exposure using all applicable radiation-control measures.18,22 In addition to wearing protective apparel, radiologic technologists and radiologists should faithfully use time, distance, and shielding principles and adequately collimate the radiographic beam.22 Some x-ray procedures increase the radiologic technologist's risk of exposure due to scatter radiation. These procedures include mobile examinations, general fluoroscopy, C-arm fluoroscopy, and interventional procedures that employ high-level control fluoroscopy.

Conclusion

Gonadal shields were first used and enacted into the federal code to protect patients from radiation-induced hereditary effects. However, studies have not linked low-dose radiation to such effects. This—combined with the enhanced detector and generator technology, the understanding that lead shielding outside the FOV affords insignificant protection to the patient, and the possibility of shielding obscuring anatomy or AEC devices—has led to this paradigm shift. Furthermore, changes in tissue-weighting factors, variation in gonadal location, and recognition that most gonadal exposure comes from internal scatter, reinforces that the routine practice of gonadal shielding for abdominopelvic imaging procedures is outdated. Radiologic technologists must use other dose reduction measures to support the ALARA principle.

About the Author

Kelli Welch Haynes, EdD, RT(R), FASRT, FAEIRS, is a professor and the Director of the School of Allied Health at Northwestern State University in Shreveport, LA. She is a member of the CARES committee and coauthor of the textbook, Radiation Protection in Medical Radiography.

References

American Association of Physicists in Medicine. AAPM Position Statement on the Use of Patient Gonadal and Fetal Shielding. 2019. Accessed June 19, 2023. https://www.aapm.org/org/policies/details.asp?id=468&type=PP

Marsh RM, Silosky M. Patient shielding in diagnostic imaging: discontinuing a legacy practice. AJR Am J Roentgenol. 2019;212:755-757. doi:10.2214/AJR.18.20508

National Council on Radiation Protection. NCRP recommendations for ending routine gonadal shielding during abdominal and pelvic radiography. NCRP Statement No. 13. January 12, 2021. Accessed June 19, 2023. https://ncrponline.org/wp-content/themes/ncrp/PDFs/Statement13.pdf

Recommendation for the use of specific area gonad shielding on patients during medical diagnostic x-ray procedures. 21 CFR § 1000.50 (1976). Accessed June 27, 2023. https://www.govinfo.gov/app/details/CFR-2022-title21-vol8/CFR-2022-title21-vol8-sec1000-50

Kowalczyk N. Update on AAPM Position Statement on the use of patient gonadal and fetal shielding. Radiol Sci Edu. 2020;21:3-10.

Fauber TL. Gonadal shielding in radiography: a best practice? Radiol Technol. 2016;88:127-134.

Shore RE, Beck HL, Boice JD, Jr, et al. Recent epidemiological studies and the linear no-threshold model for radiation protection: considerations regarding NCRP Commentary 27. Health Phys. 2019;116:235-246. doi:10.1097/HP.0000000000001015

Bushong SC. Radiologic Science for Technologists: Physics, Biology, and Protection. 12th ed. Elsevier; 2021.

Feng L, Qin L, Guo D, et al. Immunological mechanism of low-dose priming radiation resistance in walker-256 tumor model mice. Exp Ther Med. 2017;14:3868-3873. doi:10.3892/etm.2017.4975

Murphy A, Chmiel E. Fractionation (radiation therapy). Radiopaedia.org. Accessed June 19, 2023. doi:10.53347/rID-71384

International Commission on Radiological Protection. The 2007 recommendations of the International Commission on Radiological Protection. ICRP publication no. 103. Ann ICRP. 2007;37(2-4). Accessed June 19, 2023. https://www.icrp.org/publication.asp?id=ICRP%20Publication%20103

Moore QT. Manipulation of projection approach in pediatric radiography. Radiol Technol. 2015;86:481-489.

Strauss KJ, Gingold EL, Frush DP. Reconsidering the value of gonadal shielding during abdominal/pelvic radiography. J Am Coll Radiol. 2017;14:1635-1636. doi:10.1016/j.jacr.2017.06.018

Committee opinion no. 723: guidelines for diagnostic imaging during pregnancy and lactation. Obstet Gynecol. 2017;130:e210-e216. doi:10.1097/AOG.0000000000002355

McKenney S, Gingold E, Zaidi H. Gonadal shielding should be discontinued for most diagnostic imaging exams. Med Phys. 2019;46:1111-1114. doi:10.1002/mp.13409

Frantzen MJ, Robben S, Postma AA, et al. Gonad shielding in paediatric pelvic radiography: disadvantages prevail over benefit. Insights Imaging. 2012;3:23-32. doi:10.1007/s13244-011-0130-3

Kaplan SL, Strauss KJ, Magill D, et al. Quantification of increased patient radiation dose when gonadal shielding is used with automatic exposure control. J Am Coll Radiol. 2020;17:1698-1704. doi:10.1016/j.jacr.2020.05.027

Bryan WB, DeMaio D, George C, et al. Update on the elimination of fetal and gonadal shielding during abdominopelvic radiography. Radiol Technol. 2022;94:143-150.

American Registry of Radiologic Technologists. Educator update: ARRT issues statement on gondal and fetal shielding. Accessed June 19, 2023. https://www.arrt.org/pages/arrt-issues-statement-on-gonadal-and-fetal-shielding

Joint Review Committee on Education in Radiologic Technology. JRCERT Position Statement on Gonadal Shielding in the Clinical Setting. 2021. Accessed June 19, 2023. https://www.jrcert.org/wp-content/uploads/2022/12/JRCERT-Gonadal-Shielding-Position-Statement.pdf

Sherer MA, Visconti PJ, Ritenour ER, et al. Radiation Protection in Medical Radiography. 9th ed. Elsevier; 2022.