Pediatric Fluoroscopy

Jessyca B. Wagner, PhD, RT(R), CIIP

Outline

- Introduction

  - History of fluoroscopy

  - Explanation of fluoroscopy

  - Comparison of fluoroscopy to other imaging modalities

- Fluoroscopic Procedures - General

  - Gastrointestinal System

  - Biliary System

  - Genitourinary System

  - Musculoskeletal/Spinal

  - Cardiovascular/Interventional

- Fluoroscopic Procedures - Pediatric

  - Gastrointestinal

  - Genitourinary

  - Interventional

- Indications and Considerations for Pediatric Fluoroscopy

  - Indications by Procedure

  - General Pediatric Considerations

  - Radiation Safety Considerations

- Conclusion

1.5 credits

Word count: 4215 (4503 with abstract and highlights)

Learning Objectives

Upon completion of this directed reading, readers will be able to:

- EXPLAIN fluoroscopy.

- IDENTIFY the medical professionals involved in pediatric fluoroscopy.

- LIST types of fluoroscopic procedures used for pediatric patients.

- COMPARE the use of fluoroscopy in adults versus children.

- RECOGNIZE the indications for pediatric fluoroscopic examinations.

Highlights

- Fluoroscopy uses ionizing radiation to create dynamic, real-time images of the internal structures of the body. - Fluoroscopy procedures can be performed on pediatric patients; however, whenever it is possible to achieve the same imaging goal with non-ionizing modalities, those options should be used first. - Radiologic technologists must consider pediatric-specific needs and challenges, such as the potential need for immobilization and/or sedation, increased radiosensitivity, and need to engage with the parent/caregiver, when performing fluoroscopy procedures on this at-risk population.

Abstract

Fluoroscopy uses ionizing radiation to create dynamic, real-time images of the internal structures of the body. This modality is used by radiologic technologists and radiologists to aid in the detection and diagnosis of disease and injury, as well as to guide medical procedures. Fluoroscopy can be used on patients of any age and status to assess and evaluate the gastrointestinal, biliary, genitourinary, musculoskeletal, neurologic, and cardiovascular systems. It is important for radiologic technologists and radiologists to follow the ALARA (as low as reasonably achievable) principle and adhere to time, distance, and shielding protocols to limit radiation dose to themselves and their patients. Medical imaging professionals should wear lead aprons to limit occupational exposure while using proper technique, centering, positioning, and collimation for the patient. This is especially important for the pediatric patient population, as they are more radiosensitive than adult patients. Most fluoroscopy procedures that can be performed on adult patients can also be used with pediatric patients; however, if it is possible to use ultrasound or magnetic resonance imaging to achieve the same goal, those non-ionizing modalities should be the first choice. This article will discuss the history of fluoroscopy, common procedures and indications for adult and pediatric patients, and special considerations for the pediatric population.

## History of Fluoroscopy

Fluoroscopy is a form of medical imaging that demonstrates movement and functionality of the internal structures of the body in real time using ionizing radiation, often with the addition of contrast media.1-3 This modality is also used to guide medical procedures both in and out of the medical imaging department.4,5 Fluoroscopy machines can be permanently mounted in the department or mobile, as a C-arm that can be taken into the operating room, patient rooms, and the emergency department.

The origin of fluoroscopy can be traced back to 1896, one year after Wilhelm Roentgen discovered x-rays. Thomas Edison developed the first fluoroscope, which he coined the Edison Vitascope, just 4 months after the discovery of x-rays.6-11 This device consisted of only a cardboard tube coated in fluorescent metal salt held over an x-ray tube.8,10,11 After testing over 8000 materials, Edison determined that calcium tungstate fluoresced the brightest, and that is what he used in his fluoroscope.6,7 However, the image was very dim, so the viewer had to “dark-adapt” by sitting in a darkened room for a prolonged period or wear specialized viewing goggles.12 In 1950, the intensifying screen was invented and eliminated the need for the viewer to dark-adapt or wear special goggles, improved the visibility of details in the image, and allowed for indirect viewing of the image when coupled with mirror optics, which provided a significant reduction in radiation dose to the operator.12 Later, the mirror optics were changed to television screens.

Also in 1896, British physicist John Hall-Edwards made a significant advancement when he introduced the use of x-rays in surgery and later developed a specialized piece of equipment for use in surgery.9 Then in 1901, Edison’s assistant, Clarence Dally, attempted to use a fluoroscope to help locate a bullet in President McKinley’s abdomen after he was shot; however, doctors declared President McKinley too unstable to undergo the fluoroscopy, and he died a week later.11

Over the coming years, research into the field of x-rays and fluoroscopy continued and unfortunately caused many injuries and deaths among those who studied it, especially among Edison’s team.6,7,11 Four years after developing the fluoroscope, Edison’s stopped pursuing the use of radiation when he started losing his eyesight.11 Around the same time, his assistant, Clarence Dally, developed skin lesions on his face and hand; lost his hair, eyebrows, and eyelashes; and noticed heavy wrinkles in his face, along with intense pain and swelling in his left hand—which he had always used in the fluoroscope as his imaging test subject.11 When this happened, instead of stopping his pursuit of further knowledge, Dally simply switched hands.11 He assumed that with time away from the tubes and rest, he would heal. However, by 1902, the pain in Dally’s hands had intensified and they appeared scalded. Despite many skin grafts from his legs, the lesions on his hands remained. Doctors finally determined that he had developed skin carcinoma on his left hand, and Dally agreed to have the arm amputated.11 Less than 1 year later, doctors had to remove 4 fingers from his right hand. In 1903, Dally had to have his right arm amputated, and by 1904, at age 39 years, his body was so ravaged with metastatic skin cancer that he died after 8 years of experimenting with radiation.7,11

The loss of Dally’s life, while tragic, led to awareness of the dangers of working with radiation.7,11 However, many still debated the origin of many of the injuries and deaths in those who worked with x-rays, so it was many years before proper regulations were put in place.6 In the 1930s, many shoe stores were equipped with a shoe-fitting fluoroscope, which allowed clerks and customers to examine the effect of a shoe on the bones of the feet. After more information came to light regarding the effects of unnecessary radiation doses, these devices were banned from use in the United States in the 1960s but remained in use in parts of Britain and Canada for another decade.6

## How Fluoroscopy Works

Fluoroscopy works much like traditional diagnostic radiography in that an x-ray beam passes through the patient, strikes an image detector, and produces an image. However, there are a few additional steps with fluoroscopy that allow for the image to be dynamic instead of static. Fluoroscopy machines can be configured in 2 ways: with an image intensifier or with a digital flat-panel detector. Both methods offer real-time, dynamic images of the patient’s body.10,12 In both configurations, the x-ray tube is mounted under the x-ray table and sends the radiation up through the patient. This is the opposite of traditional diagnostic radiography.10,12 In units with an image intensifier, the fluoroscopy imaging chain starts with the x-ray tube and collimator, which sends the beam of radiation up through the radiolucent table into the patient. Here, the x-ray photons interact with the patient and are absorbed, transmitted, or scattered. The photons that pass through the patient strike a grid attached to the image intensifier. The image intensifier sends the beam through an optical coupling device and video camera, where it is digitized and sent to a television monitor for viewing by the radiologist during the examination.10,12 The process is the same in units with a flat-panel detector, until the beam passes through the patient. At that point, the photons strike a grid attached to the flat-panel detector and the detector converts the analog image to a digital format, sending the information to a computer where it can be visualized on a monitor by the radiologist.10,12 Using a flat-panel detector for fluoroscopy is more efficient and results in a higher quality image that can be produced with less radiation, resulting in a lower dose for the patient and the medical imaging professionals performing the examination.9,10,12

Common Uses for Fluoroscopy and Training Standards

Fluoroscopy has a variety of applications and uses in the medical field to evaluate many systems in the human body. It can be used alone as its own examination or with other diagnostic and therapeutic procedures.9,10 Most fluoroscopy examinations are completed by a radiologist or radiologist assistant and assisted by a radiologic technologist. Occasionally, a radiology nurse may also be included depending on the examination.1,2 All personnel involved in fluoroscopy procedures are expected to have a certain level of training and qualifications to perform all examinations safely. The physician performing the examination must understand anatomy, physiology, and pathophysiology, as well as have the ability to perform clinical and imaging evaluation of patients to ensure the proper procedure is being performed. They also must meet all educational and competency requirements outlined by the American College of Radiology (ACR) to perform or supervise fluoroscopy procedures.5 The physician is assisted by the radiologist-led team, which may include a non-physician radiology provider such as a registered radiologist assistant, a radiologic technologist, and other ancillary personnel as needed. These providers also must meet educational and competency requirements outlined by the ACR to ensure all examinations performed are of the highest quality and exhibit exemplary patient care.5

Common uses for fluoroscopy include, but are not limited to, imaging of the gastrointestinal (GI) and biliary systems,1,2,8-10,13 genitourinary imaging,3,10,13 musculoskeletal imaging,1,2,8-10 and cardiovascular and interventional radiology.1,2,9,10,14 Examples of specific uses are included in Table 1.1-3,8-10,13,14

| Table 1. Common Uses of Fluoroscopy |  |
| --- | --- |
| GI and Biliary1,2,8-10,13 | - Barium studies of the GI tract from the esophagus to the colon - Placement and assessment of various GI tubes and catheters - Examination of the biliary system - Retrieval of biliary stones - Placement of catheters in the bile duct in interventional procedures |
| Genitourinary3,10,13 | - Examination of the female reproductive organs - Examination, catheterization, and stenting of the urinary system from the kidneys to the bladder - Retrieval and/or lithotripsy of urinary stones |
| Musculoskeletal1,2,8-10 | - Evaluation of bones and joints - Arthrography and myelography to aid in the injection of contrast media into joint spaces or vertebral spaces to better visualize injuries with cross-sectional modalities such as MRI or CT - Lumbar punctures - Assisting in orthopedic surgeries (eg, fracture repairs, spinal fusions, and joint replacements) |
| Cardiovascular and Interventional Radiology1,2,9,10,14 | - Evaluation of the heart and vessels with angiography, arteriography, or venography - Intervention of cardiac problems with angioplasty, cardiac stenting, embolization, and IVC filters - Placement and assessment of ports, central lines, PICC lines, and other IV catheters - Image-guided needle biopsies - Drainage and aspiration of abscesses and fluid collections - Feeding tube placement - Cancer treatment |

CT: computed tomography; GI: gastrointestinal; IV: intravenous; IVC: inferior vena cava; MRI: magnetic resonance imaging; PICC: peripherally inserted central catheter.

## Comparison of Fluoroscopy With Other Imaging Modalities

While fluoroscopy use has been declining in recent years in favor of ultrasound or cross-sectional imaging modalities such as magnetic resonance imaging (MRI) and computed tomography (CT), its unique ability to offer the radiologist a real-time view inside the body cannot be dismissed.15 Traditional diagnostic radiography only offers static images of the body and modalities such as MRI, CT, and nuclear medicine do not allow repositioning of the patient during the examination. However, with fluoroscopy, the physician can see what is happening inside the patient in real-time, similar to a movie, with dynamic images and they can reposition the patient during the examination to get a better look at any abnormality that is observed.15 Fluoroscopy also uses less radiation than CT. However, fluoroscopy requires radiologists and radiologic technologists to be proficiently trained in a hands-on environment, and with reimbursement for fluoroscopic procedures declining, medical imaging professionals are often not given the opportunity to receive the training that is needed, leading to frustration and subpar images.15

# Fluoroscopic Procedures Commonly Used in Pediatric Patients

The most common pediatric fluoroscopic procedures include the evaluation of swallowing and the GI tract using a barium swallow (esophagram), upper GI studies with or without small bowel follow-through, and single or double contrast barium enema; evaluation of the urinary tract with an intravenous pyelogram/intravenous urogram, retrograde urogram, and voiding cystourethrogram; and guidance of interventional procedures.10 A more detailed list of the most common fluoroscopy procedures is provided in Table 2,14,16-34 along with a brief description of each examination. The processes for performing these procedures in pediatric patients are typically very similar to those for adult patients; however, pediatric patients often require additional measures. This may include immobilization, changes in preparation, changes in contrast dosage, changes in technical factors, and bringing in additional personnel to complete the examination quickly and efficiently.

| ## Table 2. Descriptions of the Most Common Fluoroscopic Procedures |  |
| --- | --- |
| ## Fluoroscopic Procedure | ## Description |
| ## GI |  |
| Barium swallow (esophagram)16 ## Single or double contrast | ## A contrast-enhanced radiographic study commonly used to assess structural characteristics of the entire esophagus |
| ## Modified barium swallow17 Performed with a speech pathologist | ## A contrast-enhanced radiographic study commonly used to evaluate a patient’s ability to swallow safely and effectively using items with different thicknesses and textures |
| Upper GI series18 ## Single or double contrast | ## A contrast-enhanced radiographic study commonly used to assess structural characteristics of the esophagus, stomach, and small intestine |
| Small bowel follow-through19 ## Can be performed after an upper GI series or as a stand-alone examination | ## A contrast-enhanced radiographic study commonly used to assess structural characteristics of the small intestine |
| Barium enema20 ## Single or double contrast | ## A contrast-enhanced radiographic study commonly used to assess structural characteristics of the large intestine (colon) |
| Biliary |  |
| T-tube cholangiogram21 | ## An intraoperative or postoperative radiographic study used to assess biliary patency and to evaluate for a possible biliary obstruction, leak, or stone using water-soluble contrast |
| Endoscopic retrograde cholangiopancreatography22 | ## A procedure that combines upper GI endoscopy and contrast-enhanced fluoroscopy to treat problems of the bile and pancreatic ducts |
| ## Genitourinary |  |
| Cystogram23 ## May be performed in radiology or the OR | ## A contrast-enhanced radiographic study commonly used to assess structural characteristics of the bladder |
| ## Voiding cystourethrogram23 | ## A contrast-enhanced radiographic study that is commonly used to assess structural characteristics of the bladder and to evaluate possible vesicoureteral reflux. The examination is similar to a cystogram, except the patient then voids their bladder under fluoroscopy |
| ## Retrograde urethrogram24 | ## A contrast-enhanced radiographic procedure performed most commonly in male patients to diagnose urethral pathology, such as trauma to the urethra or urethral stricture |
| ## IV pyelogram/IV urogram25 | ## A contrast-enhanced radiographic study commonly used to assess the structural characteristics of the kidneys, ureters, and bladder |
| ## Hysterosalpingogram26 | ## A contrast-enhanced radiographic study performed in female patients to evaluate the shape of the uterus, to determine if the fallopian tubes are open, and to investigate miscarriages because of problems in the uterus |
| ## Musculoskeletal |  |
| Arthrogram27,28 Can be performed on any joint; combined with either MRI or CT | ## Fluoroscopy-guided injection of contrast media into a joint space to enhance the visibility of injury to the tendons, ligaments, and cartilage; typically followed by MRI or CT |
| Joint aspirations27 ## Can be performed for any joint | ## Fluoroscopy-guided removal of fluid from a joint space, especially with suspected infection; can be performed with an arthrogram |
| Joint injections29 ## Can be performed in any joint, spine, or SI joints | ## Fluoroscopy-guided injection of corticosteroid with or without local anesthetic into joint space, spine facets, or SI joints for pain relief |
| ## Lumbar puncture30 | ## Fluoroscopy-guided removal of a small amount of cerebrospinal fluid or injection of medication or other substance into the lumbar region of the spinal canal |
| Myelogram28,31 ## Cervical, thoracic, or lumbar | ## Fluoroscopy-guided injection of contrast media into the thecal sac of the spinal canal to examine the spinal cord and assess spinal and disc problems; typically followed by CT when MRI is contraindicated |
| ## Cardiovascular and interventional radiology |  |
| ## Angiography28 | ## X-ray of the arteries and veins to find a blockage or narrowing of the vessels, as well as other problems |
| ## Angioplasty32 | ## Minimally invasive treatment in which a catheter is used to open a blocked blood vessel to improve the blood flow |
| ## Cancer treatment14 | ## Targeted injection of cancer treatment medication (such as chemotherapy) directly into the tumor |
| ## Cardiac catheterization33 | ## Procedure in which a thin, flexible tube (catheter) is guided through a blood vessel to the heart to diagnose or treat certain heart conditions, such as clogged arteries or irregular heartbeats |
| ## Central venous access32 | ## Insertion of a tube beneath the skin and into the blood vessels so patients may receive medication or nutrients directly into the blood stream |
| ## Embolization32 | ## Procedure in which medical materials or devices are placed into a blood vessel to temporarily or permanently stop bleeding |
| ## Image-guided biopsy14 | ## Insertion of a small needle into almost any part of the body, guided by imaging techniques, to take a tissue biopsy |
| ## IVC filter placement32 | ## Insertion of a preventive tool inside the IVC to trap any blood clots that may break free from the veins in the legs and potentially reach the heart or the lungs |
| ## Percutaneous drainage32 | ## Insertion of a catheter through a small nick in the skin to the site of an infection or abscess to collect and drain the fluid |
| ## Stent placement32 | ## Placement of a woven tube covered by a metal mesh at the site of weakness to reinforce the arterial wall |
| ## Thrombolysis32 | ## Treats vascular blockages and improves blood flow by dissolving abnormal blood clots |
| Tumor ablation32 ## Microwave or radiofrequency | Microwave ablation uses electromagnetic waves to destroy a tumor; radiofrequency ablation uses radio waves to create heat and damage tissue to decrease pain by damaging pain-sensing nerves or to treat cancer by damaging tumor cells |
| Vertebroplasty/kyphoplasty32,34 | Fluoroscopy-guided treatment of fractured vertebrae by injecting a cement mixture into the fractured bone through a hollow needle with (kyphoplasty) or without (vertebroplasty) the use of a balloon |

CT: computed tomography; GI: gastrointestinal; IV: intravenous; IVC: inferior vena cava; MRI: magnetic resonance imaging; SI: sacroiliac.

# Indications for Using Fluoroscopy in Pediatric Patients by Procedure

While most indications for pediatric fluoroscopy examinations are the same as in adult patients, there are some indications that are unique to pediatric patients. Common indications for pediatric barium swallow studies include dysphagia, odynophagia, noncardiac chest pain, recurrent pneumonia or chronic tracheobronchial inflammation, possible esophageal stricture or esophageal foreign body, or unexplained pneumomediastinum.35,36 For pediatric upper GI studies (with or without small bowel follow-through), common indications include vomiting, abdominal pain, weight loss or failure to thrive, congenital syndromes or anomalies associated with intestinal malrotation, chronic respiratory disease (including cough), preoperative evaluation prior to gastrostomy tube placement, postoperative evaluation (such as to confirm there is no leak or obstruction), and small bowel obstruction.35,36 Common indications for pediatric barium enema studies include intussusception reduction, neonatal distal bowel obstruction, to evaluate the position of the cecum in a newborn with equivocal upper GI study and potential malrotation, to evaluate for stricture after necrotizing enterocolitis, presurgical planning and possible Hirschsprung disease in infants or older children.36

Although less common in pediatric patients, an intravenous pyelogram/intravenous urogram small bowel follow-through may be performed to evaluate ureter course and integrity.36 Retrograde urethrograms, while also less common in pediatric patients, can evaluate urethral injury or stricture.36 The most common genitourinary study performed in pediatric patients is the voiding cystourethrogram. Common indications for this examination include but are not limited to prenatal hydronephrosis and urinary tract infection, pelvic fracture with a possible bladder injury, evaluation of neurogenic bladder, ambiguous genitalia, and fistula evaluation in boys with a high imperforate anus.36 Fluoroscopy can also be used in pediatric patients to check the placement of a gastrostomy tube, port, or central line or to guide interventional procedures such as those already listed above.36 Most interventional procedures require little to no change for pediatric patients other than to accommodate their smaller size.37 Indications for pediatric interventional procedures are numerous and can vary from patient to patient. The sheer number of procedures and indications goes beyond the scope of this article.

## General Considerations for Pediatric Imaging

In 2016, the US National Council on Radiation Protection and Measurements released a report stating that 16% of radiation dose to children comes from medical imaging procedures, with 3% coming from fluoroscopy.38 This is an increase from 2009, in which the estimated total medical radiation dose for children was 9%.38

Pediatric patients are not just small adults. Patients in this population have unique characteristics and needs that should be addressed by those caring for them. For several reasons, medical imaging cannot be deemed “one-size-fits-all”, and health care providers who work in the radiology department must be aware of the unique aspects of pediatric imaging.38,39 Table 3 highlights some elements that are specific to pediatric patients, which radiologic technologists, radiology nurses, radiologists, and other medical imaging professionals should consider.38,40

Table 3. Considerations for Fluoroscopy Examinations in Pediatric Patients

- Medical decisions are made by caregivers and medical professionals

- There are different clinical manifestations and imaging appearances for similar disorders in children

- Some patients cannot talk so the health care provider needs to engage with the caregiver

- It is beneficial to have a child-friendly environment and team members who are comfortable with and trained in pediatric care

- Many examinations require sedation, anesthesia, or other strategies to immobilize the patient

- There is a wide variation in patient size, anatomy, and physiology when imaging children

- Because most pediatric imaging in the United States is not performed in pediatric centers, there is the potential for inadequate configurations of ionizing radiation equipment for children

- Challenges with dose estimation, radiation sensitivity, and risk estimation are present when imaging children

Proper training in pediatric care is imperative for all health care providers, especially those professionals who work in a nonpediatric facility.38-40 Radiologic technologists and other personnel must remember that pediatric patients often cannot communicate what is wrong or may not understand what is happening or why they are in the hospital, which can be scary. This is particularly true in the medical imaging department because of the large machines and unfamiliar environment. Health care providers in the medical imaging department also need to be trained on ways to interact with pediatric patients to earn their trust and cooperation and ensure quality images are obtained during the examination, and radiologists must be trained to correctly interpret the images that are taken given the wide variation in the size, anatomy, and physiology of pediatric patients.40

Often, medical imaging examinations are long and the patient needs to remain motionless for the duration of the examination. This can be difficult for many patients, especially those in the pediatric population. For these patients, it may be necessary to immobilize or restrain them. Options for immobilizing pediatric patients include swaddling the patient, distracting them, verbal techniques (eg, xx, xx), having the parent or radiologic technologist physically hold the patient, and using devices such as the Pigg-O-Stat, Velcro straps, and sand bags.41 Sedation also may be necessary for some procedures, such as MRI.40,41 It is imperative for radiologic technologists and other health care providers to obtain consent from the patient, if possible, before using any immobilization techniques. In this instance, even if consent is given by the parent or caregiver, failure to obtain consent from the pediatric patient changes immobilization to restraint, which can increase the patient’s anxiety and result in post-traumatic stress disorder for future imaging procedures.41

## Radiation Safety Considerations

Radiation safety should be a primary concern for radiologic technologists and radiologists, regardless of the age of the patient; however, it is especially important when imaging pediatric patients because they are more radiosensitive.37-40 Children are 2 to 5 time (some even cite 10 times) more sensitive to radiation exposure than adults because of the large number of rapidly dividing cells in their bodies and their longer life expectancy, which increases the likelihood of long-term effects.38-42

Radiation doses can vary depending on the modality being used, the method of dose measurement, the age and size of the patient, the examination being performed, and they technical factors used in the examination.38,39 The dose of fluoroscopy procedures is typically described using dose rates because the dose depends on the imaging time and number of images taken.39 Ionizing radiation can result in either direct or indirect damage to DNA. Single-stranded DNA damage can be repaired easily; however, double-stranded DNA damage is difficult to repair completely.39 Biological effects range from immediate cell necrosis, which can lead to organism death, to cell damage, which causes prolonged effects such as birth defects and cancer (the most likely prolonged effect to result from medical imaging exposure).39 Radiation exposure can cause 2 types of biological effects: deterministic and stochastic. Deterministic effects occur at a threshold dose, typically beyond that used for diagnostic imaging, and can include epilation, skin burns, and cataracts. Stochastic effects do not have a threshold dose and instead are believed to increase in risk and severity as the radiation dose increases. These effects include radiation carcinogenesis and radiation-induced genetic damage.39

Before performing a fluoroscopy examination or other medical imaging examination with the potential for elevated doses of radiation, health care providers should check the patient’s medical record for related recent studies and note the date, modality, findings, and location of the images to ensure the procedure is medically necessary. Once the examination has been performed, all available radiation dose data should be recorded in the patient’s medical record.5 To limit radiation dose, the as low as reasonably achievable (ALARA) principle should always be used, along with time, distance, and shielding.35,39,40,43 Radiation exposure always should be limited to as little time as possible; health care providers should distance themselves as far from the primary beam as possible; and appropriate technical factors, proper centering, and collimation should be used at all times. Health care providers should also shield themselves to limit occupational exposure. The guidelines on patient shielding have recently changed and patient gonadal and fetal shielding are no longer recommended.44 Communicating with the patients and their caregivers is important, not only to aid in cooperation and trust, but also to inform them of the risks versus benefits of the examination that is being performed.

There have been worldwide efforts to increase education about pediatric medical imaging for both consumers and health care providers through organizations such as the National Cancer Institute, the Health Physics Society, the American Academy of Pediatrics, and the American College of Radiology, as well as efforts to regulate pediatric radiation dose.39 In the United States, one major organization is the Image Gently® Alliance, which is a coalition of health care providers whose goal is to change practice to provide “safe, high-quality pediatric imaging worldwide.”39,45 Image Gently® has a Pause and Pulse campaign related to fluoroscopy that reminds providers to pause and child-size the technique, use the lowest pulse rate possible, and consider ultrasound or MRI when possible. This suggested change in modality is one primary difference between adults and children for fluoroscopy.40,46 While fluoroscopy is often recommended in adult patients over a modality with more radiation, such as CT, non-ionizing modalities such as ultrasound and MRI are recommended for pediatric patients to eliminate dose altogether when possible.40

# Conclusions

Fluoroscopy is a useful medical imaging modality to aid in the detection and diagnosis of numerous diseases and injuries and to guide life-saving medical procedures. This modality is versatile and can be used in patients of all ages to evaluate many systems of the body, including the gastrointestinal, biliary, genitourinary, musculoskeletal, neurologic (spinal), and cardiovascular systems. However, care must be taken when this—or any medical imaging procedure that uses ionizing radiation—is performed on a pediatric patient. When possible, ultrasound and MRI should be the modalities of choice in pediatric patients to limit or eliminate applying any radiation dose to their more radiosensitive tissues and lessen the potential side effects. When fluoroscopy and ionizing radiation are needed, radiologic technologists and radiologists should always practice the ALARA principle and remember the rules for time, distance, and shielding to protect themselves and the patient from unnecessary radiation exposure.

References

1. US Food and Drug Administration. Fluoroscopy. Updated February 21, 2023. Accessed May 23, 2023. https://www.fda.gov/radiation-emitting-products/medical-x-ray-imaging/fluoroscopy

2. US Centers for Disease Control and Prevention. Radiation in healthcare: fluoroscopy. Reviewed October 20, 2021. Accessed May 23, 2023. https://www.cdc.gov/nceh/radiation/fluoroscopy.html

3. US National Library of Medicine. Fluoroscopy. Medline Plus. Updated November 12, 2021. Accessed May 23, 2023. https://medlineplus.gov/lab-tests/fluoroscopy/

4. Butler PF. Fluoroscopy protocols. Image Wisely. Accessed May 23, 2023. https://www.imagewisely.org/Imaging-Modalities/Fluoroscopy/Protocols

5. American College of Radiology. ACR–AAPM technical standard for management of the use of radiation in fluoroscopic procedures. 2018. Accessed May 23, 2023. https://www.acr.org/-/media/ACR/Files/Practice-Parameters/MgmtFluoroProc.pdf

6. University of Washington Medicine. Featured history: fluoroscopy. Department of Radiology. Accessed May 23, 2023. https://rad.washington.edu/blog/featured-history-fluoroscopy/

7. Kansas State University. Physics and the detection of medical x-rays: fluoroscope. Physik im Kontext (PIKO). Accessed May 23, 2023. https://web.phys.ksu.edu/mmmm/piko/fluor.htm#:~:text=The%20fluoroscope%20was%20invented%20in,hazard%20related%20to%20the%20fluoroscope

8. Tharaga, Thushya, Vishnuga. Fluoroscopy: History. Idaho State University. Accessed May 23, 2023. https://biologyisu.wordpress.com/history-2/

9. Kingu M. Fluoroscopy: a history of advancements and applications. LinkedIn. April 14, 2023. Accessed May 23, 2023. https://www.linkedin.com/pulse/fluoroscopy-history-advancements-applications-mkoma-kingu-/

10. Gearing L. Pediatric fluoroscopy: clinical and practical care and management. eRadImaging. June 15, 2022. Accessed May 23, 2023. https://www.eradimaging.com/course/985

11. King G. Clarence Dally – the man who gave Thomas Edison x-ray vision. Smithsonian Magazine. March 14, 2012. Accessed May 23, 2023. https://www.smithsonianmag.com/history/clarence-dally-the-man-who-gave-thomas-edison-x-ray-vision-123713565/

12. Johnston JN, Fauber TL. Fluoroscopic imaging. In: Essentials of Radiographic Physics and Imaging. 3rd edition. Elsevier; 2020:216-236.

13. Houston JD, Davis M. Fundamentals of Fluoroscopy. W. B. Saunders Company; 2001. Accessed May 23, 2023. http://www.osumcradiology.org/resources/Rotation-Books/Fundamentals-of-Fluoroscopy.pdf

14. Johns Hopkins Medicine. Interventional radiology. Accessed May 23, 2023. www.hopkinsmedicine.org/health/treatment-tests-and-therapies/interventional-radiology

15. Shalom NE, Gong GX, Auster M. Fluoroscopy: an essential diagnostic modality in the age of high-resolution cross-sectional imaging. World J Radiol. 2020;12:213-230. doi:10.4329/wjr.v12.i10.213

16. Chen A, Tafti D, Tuma F. Barium swallow. In: StatPearls [Internet]. StatPearls Publishing; 2023. Accessed May 23, 2023. https://www.ncbi.nlm.nih.gov/books/NBK493176/

17. RadiologyInfo.org for Patients. Video fluoroscopic swallowing exam (VSFE)/esophagram. Reviewed June 1, 2022. Accessed May 24, 2023. https://www.radiologyinfo.org/en/info/modbariumswallow

18. RadiologyInfo.org for Patients. Upper GI x-ray. Reviewed April 15, 2022. Accessed May 24, 2023. https://www.radiologyinfo.org/en/info/uppergi

19. RadiologyInfo.org for Patients. Small bowel follow-through. Reviewed April 15, 2022. Accessed May 24, 2023. https://www.radiologyinfo.org/en/info/small-bowel-follow-thru

20. RadiologyInfo.org for Patients. Lower GI x-ray. Reviewed May 1, 2023. Accessed May 24, 2023. https://www.radiologyinfo.org/en/info/lowergi

21. UT Southwestern Department of Radiology. T-tube cholangiogram. Reviewed June 4, 2015. Accessed May 25, 2023. https://www.utsouthwestern.edu/education/medical-school/departments/radiology/protocols/assets/T-Tube%20Cholangiogram.pdf

22. National Institute of Diabetes and Digestive and Kidney Diseases. Endoscopic retrograde cholangiopancreatography (ERCP). Reviewed June 2016. Accessed May 25, 2023. https://www.niddk.nih.gov/health-information/diagnostic-tests/endoscopic-retrograde-cholangiopancreatography

23. Johns Hopkins Medicine. Cystography. Accessed May 25, 2023. https://www.hopkinsmedicine.org/health/treatment-tests-and-therapies/cystography

24. Burks FN, Santucci RA. Urethrogram. Medscape. Updated September 24, 2018. Accessed May 25, 2023. https://emedicine.medscape.com/article/1893948-overview

25. RadiologyInfo.org for Patients. Intravenous pyelogram (IVP). Reviewed April 15, 2022. Accessed May 25, 2023. https://www.radiologyinfo.org/en/info/ivp

26. RadiologyInfo.org for Patients. Hysterosalpingography. Reviewed June 1, 2022. Accessed May 25, 2023. https://www.radiologyinfo.org/en/info/hysterosalp

27. Murphy A. Arthrogram. Radiopaedia. Revised March 23, 2023. Accessed May 25, 2023. https://radiopaedia.org/articles/arthrogram-1?lang=us

28. Inside Radiology. Tests. Revised March 16, 2018. Accessed May 23, 2023. https://www.insideradiology.com.au/for-health-professionals/tests/

29. RadiologyInfo.org for Patients. Facet joint block. Reviewed May 1, 2023. Accessed May 25, 2023. https://www.radiologyinfo.org/en/info/facet-joint-block

30. RadiologyInfo.org for Patients. Lumbar puncture. Reviewed April 15, 2022. Accessed May 25, 2023. https://www.radiologyinfo.org/en/info/spinaltap

31. RadiologyInfo.org for Patients. Myelography. Reviewed May 1, 2023. Accessed May 25, 2023. https://www.radiologyinfo.org/en/info/myelography

32. Society of Interventional Radiology. Glossary of interventional radiology treatments. Accessed May 25, 2023. https://www.sirweb.org/patient-center/conditions-and-treatments/glossary-of-ir-treatments/

33. Mayo Clinic. Cardiac catheterization. Updated October 15, 2021. Accessed May 26, 2023. https://www.mayoclinic.org/tests-procedures/cardiac-catheterization/about/pac-20384695

34. RadiologyInfo.org for Patients. Vertebroplasty & kyphoplasty. Reviewed May 1, 2023. Accessed May 26, 2023. https://www.radiologyinfo.org/en/info/vertebro

35. American College of Radiology. ACR–SPR practice parameter for the performance of contrast esophagrams and upper gastrointestinal examinations in infants and children. Revised 2020. Accessed May 26, 2023. https://www.acr.org/-/media/ACR/Files/Practice-Parameters/UpperGI-Infants.pdf

36. University of North Carolina Radiology. Pediatric fluoro handbook. December 18, 2020. Updated January 1, 2021. Accessed May 26, 2023. https://rads.web.unc.edu/wp-content/uploads/sites/12234/2018/07/Peds-Fluoro-Manual-for-Residents-2020-v2.pdf

37. American Society of Radiologic Technologists and Image Gently®. Enhancing radiation protection during pediatric fluoroscopy. 2013. Accessed May 26, 2023. https://www.imagegently.org/Portals/6/Procedures/Module%201%20transcript.pdf

38. Frush DP, Sorantin E. Radiation use in diagnostic imaging in children: approaching the value of the pediatric radiology community. Pediatr Radiol. 2021;51:532-543. doi:10.1007/s00247-020-04924-6

39. Frush DP, Applegate KE. Radiation risk from medical imaging in children. In: Medina LS, Applegate KE, Blackmore CC, eds. Evidence-Based Imaging in Pediatrics: Optimizing Imaging in Pediatric Patient Care. Springer; 2010:25-39. doi:10.1007/978-1-4419-0922-0_3

40. Thukral BB. Problems and preferences in pediatric imaging. Indian J Radiol Imaging. 2015;25:359-364. doi:10.4103/0971-3026.169466

41. Murphy A. Pediatric immobilization. Radiopaedia. Revised March 23, 2023. Accessed May 31, 2023. https://radiopaedia.org/articles/paediatric-immobilisation?lang=us

42. Ploussi A, Brountzos E, Rammos S, et al. Radiation exposure in pediatric interventional procedures. Cardiovasc Intervent Radiol. 2021;44:857-865. doi:10.1007/s00270-020-02752-7

43. Johnson MM. Radiation protection education in fluoroscopy. Radiol Technol. 2015;86:511-528. http://www.radiologictechnology.org/content/86/5/511.abstract

44. American Association of Physicists in Medicine. AAPM Position Statement on the Use of Patient Gonadal and Fetal Shielding. 2019. PP 32-A. Accessed May 26, 2023. https://www.aapm.org/org/policies/details.asp?id=468

45. Image Gently®. The Alliance. Accessed May 26, 2023. https://www.imagegently.org/About-Us/The-Alliance

46. Image Gently®. Pause and pulse: image gently in fluoroscopy. Accessed May 26, 2023. https://www.imagegently.org/Procedures/Fluoroscopy