Introduction to Portable Magnetic Resonance Imaging

Jessyca B. Wagner, Ph.D., RT(R), CIIP

Outline:

- Introduction

- Fixed MRI

  - History of MRI

  - Explanation of fixed MRI

- Portable MRI

  - Explanation of portable MRI

  - Comparison of fixed and portable MRI

  - Uses for portable MRI

  - Advantages of portable MRI

  - Disadvantages of portable MRI

- Conclusion

Word count: TBD

Credits: 0.5

Learning Objectives:

Upon completion of this directed reading, readers will be able to:

- Explain magnetic resonance imaging (MRI).

- Compare fixed MRI and portable MRI.

- IDENTIFY uses for portable MRI.

- Discuss advantages and disadvantages of portable MRI.

Highlights

- Magnetic resonance imaging (MRI) provides detailed three-dimensional images of soft tissues of the body, such as the brain, breast, and musculoskeletal system; however, not everyone can undergo this form of imaging. - Portable magnetic resonance imaging (MRI) offers a solution for patients who are not candidates for fixed MRI exams because of the decreased cost, portability, and weaker magnetic field. - It is important to note that while there are many advantages to portable MRI, the image quality and field of view are decreased and the scan time may be increased.

Abstract

Magnetic resonance imaging (MRI) is a non-invasive, non-ionizing imaging modality that can be used to create three-dimensional images of soft tissues in the body. It is primarily used for musculoskeletal, neurological, and breast imaging; however, the high purchase and operating costs, need for large space with safety zones to house it, specialized personnel, and high cost to the patient make high-field fixed MRI less accessible to those in low- and middle-income countries, rural areas, and even patient populations such as obese patients, pediatric patients, and critically ill patients. In an effort to decrease cost and increase accessibility, low-field, portable MRI machines have recently been approved by the FDA, making this modality available to more patients than ever before. Portable MRI, also known as low-field or point-of-care MRI, allows advanced imaging to be performed bedside, in nursing homes and assisted living facilities, in a patient’s home, and even in the field. While the image quality is not as high as with fixed MRI, technological advancements and AI software developments have vastly improved the scan time, image quality, and feasibility of this scanner.

# Portable Magnetic Resonance Imaging

Magnetic resonance imaging (MRI) is an imaging modality that produces three dimensional detailed images of a patient’s anatomy, especially soft tissue areas such as the brain, breast, and musculoskeletal system. The procedure is non-invasive and can be used with or without a contrast medium for disease detection, diagnosis, and treatment monitoring. Since its inception, medical imaging professionals have sought to improve the image quality and capabilities of MRI. Until recently, the primary method of improving MRI was to increase the magnetic field strength. However, these high-field scanners are costly, as they require specialized personnel, use a large amount of electrical power and costly cryogens, and are expensive to purchase and maintain.2 These requirements may mean that MRI is not possible in smaller facilities, rural areas, or low- and middle-income countries. As an alternative, low-field MRIs have recently been introduced that are “portable, cryogen free, easy to use, provide rapid patient loading and unloading, have minimal power requirements, and have relatively low purchase prices and maintenance costs.”2 These low-field scanners have the potential to revolutionize MRI imaging because they offer point-of-care diagnostic capabilities, can be used in a variety of settings thanks to their portability and lower cost, and even have the potential of being utilized in ambulances and on the battlefield in the future.2

# Fixed MRI

## History of MRI

Nikola Tesla first described the rotating magnetic field in 1882 and in 1937, Isidor Rabi developed nuclear magnetic resonance (NMR), which earned him the 1944 Nobel Prize in Physics. While Rabi’s method was primarily used for chemical substances, in the 1960s Raymond Damadian began investigating the application of NMR on living organisms. In 1971, he determined that using radio waves and measured emissions from hydrogen atoms, he could detect cancerous tissue because it contained more water than healthy tissue.3 Around the same time, Paul Lauterbur, a chemist, was using NMR to create images of vials of water, clams, and green peppers. After Damadian published his 1971 findings, Lauterbur realized his work could be translated to biomedical applications. “He was the first to realize that a gradient magnetic field would allow observers to take two-dimensional images of an object, which could then be stacked to create a three-dimensional view – an idea he first sketched out on a paper napkin between bites of a Big Boy hamburger.”3 Simultaneously with Damadian and Lauterbur’s work, physicist Peter Mansfield was trying to complete scans in minutes rather than hours. He adopted a new technique he called “line scan imaging” and could capture images of his grad student’s finger in only 15-23 minutes per section, marking the first time human anatomy was successfully scanned using NMR technology.3

Indomitable, the first full-body scanner, was designed by Damadian. It was a 1.5-ton machine that required the “patient” to wear a cardboard-backed antenna coil and lay inside the scanner. The first attempt on Damadian himself was unsuccessful; however, his grad student volunteered to try once it was determined Damadian had no side-effects. This scan was successful, providing a two-dimensional image of the student’s chest. The team hypothesized the second scan was successful because the student was thinner and had less body fat than Damadian.3 Eventually Damadian’s methods were deemed too slow for routine use and Lauterbur and Mansfield’s methods were adopted, earning them the 2003 Nobel Prize in Medicine.3 Clinical-grade scanners became more commonplace in the 1980s after they were approved by the FDA.2

## Explanation of Fixed MRI

Unlike computed tomography (CT) or diagnostic radiography, MRI scans do not use ionizing radiation. Instead, MRIs use powerful magnets that produce a strong magnetic field, forcing protons in the body to align with the field. When a current of radio waves is pulsed through the patient, the protons are stimulated, spinning out of equilibrium, causing them to strain against the pull of the magnetic field.1 When the radiofrequency is turned off, the MRI sensors detect the energy released as the protons realign with the magnetic field. “The time it takes for the protons to realign with the magnetic field, as well as the amount of energy released, changes depending on the environment and the chemical nature of the molecules.”1 Because of the complexity of how the energy interacts with the patient, it is important for the patient to remain very still during the exam so the image does not blur. Contrast medium, typically Gadolinium, may be used intravenously before or during the MRI to increase the speed of the proton realignment. The faster the protons realign, the brighter the image.1

MRI is best suited for the soft tissues of the body. The brain, spinal cord, nerves, muscles, ligaments, and tendons are seen more clearly with MRI than with CT or diagnostic radiography, which are better suited for the bony structures of the body.1 MRI is often used to assess sports injuries, since these often involve the soft tissue of areas of the joints. MRI is also used frequently to diagnose aneurysms and tumors in the brain because of its ability to distinguish between gray matter and white matter. There are also specialized forms of MRI, such as functional MRI (fMRI) that can observe brain structures and determine which areas of the brain activate during various tasks and activities.1 The major downside to using MRI is that it is typically more expensive than CT or diagnostic radiography.

Fixed MRI is the traditional form of MRI that is housed permanently in an imaging department, so patients are required to travel to a hospital or imaging center for a scan. Fixed MRI scanners are traditionally high-field scanners, meaning their magnetic strength ranges from 1.5-3 and even up to 7 Tesla (T). The magnet is always on and pulls with a force of 30,000 – 60,000 times that of Earth’s gravity. This high magnetic strength requires a dedicated imaging suite with safety zones to protect people and equipment from the magnet. The American College of Radiology (ACR) has defined four safety zones for MRI facilities. The zones correspond to the increasing potential safety risks as the proximity to the magnet increases.5 Zone I includes all areas freely accessible to the public such as the waiting room. No screening is required in Zone I and anyone can be in this zone because it is typically located outside the MRI environment. Zone II is the area between uncontrolled Zone I and the strictly controlled Zones III and IV. Patients may move freely in this zone under MRI personnel supervision. This is typically where MRI safety screenings and patient preparation, such as dressing changes, occur. Zone III is the area near the magnet room, typically the control room, where the magnetic fields are strong enough to pose a threat to non-screened personnel or patients and ferromagnetic objects. Only screened persons are allowed in Zone III, and they must be accompanied by personnel who have been properly trained in MRI safety. Zone IV is the scanning room itself. The magnetic strength is highest in this zone and the MRI technologist must always have a clear line of sight to the scanner and any patients or personnel in the room.5

# Portable MRI

## Explanation of Portable MRI

Portable MRI, also referred to as low-field MRI or point-of-care MRI, functions the same as fixed MRI, just with a smaller magnet. Portable MRI has steadily increased in popularity in recent years. As previously mentioned, high-field fixed MRI is limited in low- and middle-income countries and even rural areas in the United States and other high-income countries because of its high cost. Because of this, approximately 90% of people lack access to MRI worldwide. Low-field MRI has advanced enough recently to become a solution with less expensive and portable devices. However, the image quality of low-field scanners is not as high as that of high-field scanners.6 The term “low-field” typically describes scanners with a magnetic field less than 1.5T but can also indicate a range as specific as 0.01-0.1T.6 While low-field scanners have existed since MRI was FDA approved in the 1980s, they were eclipsed by the high-field scanners until 2018 when renewed commercial interest led to FDA clearance of several low-field scanners thanks to technological advancements that have greatly improved the image quality.6 Hardware improvements and software developments in low-field scanners have improved the image quality obtained per unit time, allowing lower magnetic fields to produce high signal-to-noise ratio (SNR) images of diagnostic quality.6 Artificial intelligence and deep learning algorithms are also helping to denoise and improve image quality and aid in image postprocessing.6

## Comparison of Fixed and Portable MRI

Table 1 - Fixed vs Portable MRI highlights some of the common differences between fixed and portable MRI machines.6

Table 1 - Fixed vs Portable MRI

|  | Fixed MRI | Portable MRI |
| --- | --- | --- |
| Magnet Size | 1.5-7 T | Less than 1.5 T |
| Cost | $1.5 million and up | $1 million or less |
| Footprint | Permanent setup with ACR Zones Cryogenic coolers Substantial amounts of power and always on Large, heavy hardware | Less need for a dedicated shielded room No super cooling Less power and not always on Smaller and lighter hardware |
| Safety | Patients and personnel must be carefully screened Ferromagnetic objects at great risk of being pulled into the magnet Extremely loud during scanning High rate of claustrophobia | Less risk of device interactions Less risk of metallic projectiles Decreased acoustic noise Decreased claustrophobia |
| Bore Configurations | Closed bore (50-60 cm inside diameter) Open and wide-bore options available (70 cm or greater diameter) | C-shaped Wider bore Single-sided Vertical orientation |
| Image Quality | High quality High SNR Great detail | Slightly lower quality Slightly lower SNR Less detail |

## Uses for Portable MRI

Portable MRI opens up several possibilities that are impossible or difficult to attain with fixed MRI. One very common use for portable MRI is point-of-care (POC) exams in the emergency room and intensive care units. Often, patients in these areas of the hospital cannot travel to the imaging department or require interventions that are not MRI-safe (e.g., some ventilators, drug infusion pumps). The lower strength of the magnet makes it less likely to interfere with pacemaker/defibrillators or other implanted devices and less likely to draw objects into the magnet.2,6, Portable MRI is also being used in the operating room. Intraoperative MRI (iMRI) is primarily used by neurosurgeons for guidance to remove tumors and treat conditions such as epilepsy. iMRI is also used during awake craniotomies so the surgeon can monitor brain activity during the procedure.6,

Neuroimaging has been the focus of most research studies using portable MRI. Neurological disorders are the leading cause of disability and the second leading cause of death worldwide. MRI is the gold standard in neurological imaging; however, many patients cannot undergo fixed MRI scans. Patients with cardiac implantable electronic devices (pacemakers or defibrillators), metallic intraocular foreign bodies, implantable neurostimulators, cochlear implants, aneurysm clips, and other non-removable ferrous materials cannot enter an MRI suite. Portable MRI however, offers a great alternative for neuroimaging.2,4,6,7,9,,,,,,,

Pediatric MRI has increased in use for clinical and research purposes because of the lack of ionizing radiation. However, because of the cost, long scan times, need for sedation, noise sensitivity, and claustrophobia often encountered when imaging pediatric patients, its use is still limited. Portable MRI has seen an increase in patient compliance because noise sensitivity, claustrophobia, and need for sedation have all decreased. These exams can also be brought to the patient, which is often less intimidating for pediatric patients. Finally, they are cheaper, making them a more viable option for many families.2,6,11

Low-field MRI is also valuable for musculoskeletal (MSK) imaging. Orthopedics is one of the most common specialties using low-field MRI in clinical practice today. The versatility of low-field MRI, lack of interference with metal implants, and cost effectiveness make low-field MRI a popular choice in orthopedics.6

Advantages and Disadvantages of Portable MRI

There are many advantages and disadvantages to portable MRI. The primary advantages as discussed throughout this article include accessibility, portability, cost effectiveness, versatility, safety, and ease of use. Bringing MRI to underserved populations such as low- and middle-income countries, rural areas, and even critically ill, obese, and pediatric patients makes further research in low-field MRI appealing and necessary. The major disadvantages of portable MRI are decreased image quality, decreased field of view, and increased scan time.2,6 However, many researchers agree the benefits outweigh the shortcomings in many cases and the technology will continue to improve, resulting in even more widespread use of the scans.2,4,6,7,9,11,12,13,15,16,17

Conclusion

Fixed MRI has been a valuable imaging tool for medical professionals since its initial FDA approval in the early 1980s. The three-dimensional, detailed images of soft tissue provide vital information to physicians and medical imaging personnel that help with the detection, diagnosis, and treatment planning of many diseases. However, the high cost of high-field, fixed MRI as well as the housing requirements, personnel training, loud noise, and small bore keep it from being an option for all patient populations. Low-field, portable, point-of-care MRI, which was FDA approved in 2018, has given patients in low- and middle-income countries and rural areas access to MRI that was previously unattainable. Also, populations such as obese patients, pediatric patients, and critically ill patients now have an option to undergo MRI exams without needing sedation, expensive open or wide-bore machines, and without having to be transported to the imaging department. The advantages of portable MRI are great; however, the decreased image quality and increased scan time need improvement. Technological and software advancements and the integration of artificial intelligence have increased the application of portable MRI and will likely continue to improve the scanning to make it a viable option for many patients to come.

References

National Institute of Biomedical Imaging and Bioengineering (NBIB). Magnetic Resonance Imaging (MRI). National Institutes of Health. n.d. Accessed July 31, 2023. https://www.nibib.nih.gov/science-education/science-topics/magnetic-resonance-imaging-mri#main-content

Basser P. Detection of stroke by portable, low-field MRI: A milestone in medical imaging. Sci Adv. 2022;8(16):eabp9307. https://doi.org/10.1126/sciadv.abp9307

University of Washington Medicine Department of Radiology. Featured History: Magnetic Resonance Imaging. n.d. Accessed July 31, 2023. https://rad.washington.edu/blog/featured-history-magnetic-resonance-imaging/

Hovis G, Langdorf M, Dang E, et al. MRI at the bedside: A case report comparing fixed and portable magnetic resonance imaging for suspected stroke. Cureus. 2021;13(8):e16904. https://doi.org/10.7759%2Fcureus.16904

Medical Professionals. MRI Zones: A Guide for Rad Techs. October 31, 2022. Accessed July 31, 2023. https://www.medical-professionals.com/en/mri-zones-guide/

Arnold TC, Freeman CW, Litt B, et al. Low-field MRI: Clinical promise and challenges. J Magn Reason Imaging. 2023;57:25-44. https://doi.org/10.1002/jmri.28408

Cho S-M, Wilcox C, Keller S, et al. Assessing the SAfety and FEasibility of bedside portable low-field brain magnetic resonance imaging in patients on ECMO (SAFE-MRI-ECO study): Study protocol and first case series experience. Critical Care. 2022;26:119. https://doi.org/10.1186/s13054-022-03990-6

Mayo Clinic Staff. Intraoperative Magnetic Resonance Imaging (iMRI). Mayo Clinic. March 5, 2022. Accessed July 31, 2023. https://www.mayoclinic.org/tests-procedures/intraoperative-magnetic-resonance-imaging/about/pac-20394451

Cooley CZ, McDaniel PC, Stockmann JP, et al. A portable scanner for magnetic resonance imaging of the brain. Nat Biomed Eng. 2021;5(3):229-239. https://doi.org/10.1038/s41551-020-00641-5

Ghadimi M, Sapra A. Magnetic resonance imaging contraindications. Treasure Island, FL: StatPearls; 2023. https://www.ncbi.nlm.nih.gov/books/NBK551669/ Accessed August 9, 2023.

Deoni SCL, Medeiros P, Deoni AT, et al. Development of a mobile low-field MRI scanner. Sci Rep. 2022;12(1):5690. https://doi.org/10.1038/s41598-022-09760-2

Mateen FJ, Cooley CZ, Stockmann JP, et al. Low-field portable brain MRI in CNS demyelinating disease. Mult Scler Relat Disord. 2021;51:102903. https://doi.org/10.1016/j.msard.2021.102903

Mazurek MH, Cahn BA, Yuen MM, et al. Portable, bedside, low-field magnetic resonance imaging for evaluation of intracerebral hemorrhage. Nat Comm. 2021;12:5119. https://doi.org/10.1038/s41467-021-25441-6

Prabhat AM, Crawford AL, Mazurek MH, et al. Methodology for low-field, portable magnetic resonance neuroimaging at the bedside. Front Neurol. 2021;12:760321. https://doi.org/10.3389/fneur.2021.760321

Sheth KN, Mazurek MH, Yuen MM, et al. Assessment of brain injury using portable, low-field magnetic resonance imaging at the bedside of critically ill patients. JAMA Neurol. 2021;78(1):41-47. https://doi.org/10.1001/jamaneurol.2020.3263

Turpin J, Unadkat P, Thomas J, et al. Portable magnetic resonance imaging for ICU patients. Crit Care Explor. 2020;2(12):e0306. https://doi.org/10.1097%2FCCE.0000000000000306

Yuen MM, Prabhat AM, Mazurek MH, et al. Portable, low-field magnetic resonance imaging enables highly accessible and dynamic bedside evaluation of ischemic stroke. Sci Adv. 2022;8(16):eabm3952. https://doi.org/10.1126/sciadv.abm3952